# Firmware Update for Germanium Detector

The electronic system of the Germanium detector was based on a Xilinx 7z045 SoC, and detector data acquisition and processing are performed in the Programmable Logic of the SoC. A micro-SD card is used to store the firmware and related files (in BOOT sector) and a Linux system (in rootfs sector). A firmware file is a .bit file to be loaded by Linux to program the Programmable Logic of the SoC, which is done when the detector powers on. A uEnv.txt file in the BOOT sector records the name of the firmware file to be loaded.

The updating process includes adding a new firmware file and revising uEnv.txt. Both steps can be done either remotely or locally, depending on whether a valid network connection to the detector is available. The files related to the update are listed below.

## Remote update

Remote update requires a valid network connection to the detector. Depending on network configuration (check /etc/network/interfaces in the rootfs sector), when the detector powers up, it gets either a dynamic IP address from a DHCP server, or the static IP address set in this file. To check the IP address when the detector is on:

- Connect a computer to the console port (Micro USB) using a Micro USB cable.

- Open a terminal (e.g., minicom) on the computer.

- Set correct parameters for the serial port corresponding to the USB connection (baud rate 115200, 8 data bits, 1 stop bit). Open the serial port.

- Press Enter in the terminal. If it’s the first time logging in from the console port, a prompt for login appears. If seeing Linux prompt, skip to f).

- Use the provided credential to log into the Linux system of the detector.

- Run command:

$ ip a

This returns network configuration information, including the IP address of the network interface to be used (eth0).

Now the firmware file can be uploaded from the computer with command:

$ scp new-firmware.bit root@ip-address:/boot/uboot/.

Next, revise /boot/uboot/uEnv.txt. Comment the line that loads the firmware file, e.g.,

# Comment: loadxbit=fatload mmc 0 ${btaddr} top_v18b_ajk.bit; fpga loadb 0 ${btaddr} ${filesize}

Add a new line to load the new firmware file:

loadxbit=fatload mmc 0 ${btaddr} top_v20_192_384.bit; fpga loadb 0 ${btaddr} ${filesize}

Save the file and exit. Reboot the detector to load the new firmware.

$ reboot

## 2. Manual update

Manual update requires the detector to be powered off to take out the micro-SD card.

Caution: the detector shouldn’t be powered off for long time to avoid the electronic components being damaged by low temperature.

This can be done as following:

- Power off the detector.

- Remove the cover of the detector circuitry housing.

Be aware of the power cord of the cooling fan attached to the cover.

- Take out the micro-SD card.

- Insert the micro-SD card into a card reader.

- Copy the firmware file to the BOOT sector of the micro-SD card. The destination directory may differ for different systems. In Debian/Ubuntu, use command:

$ cp new-firmware.bit /media/$user/BOOT/

- Revise uEnv.txt in the BOOT sector as in the previous section.

- Unmount the micro-SD card.

- Replace the micro-SD card in the detector.

- Replace the cover. Be aware of the connection of the power cord of the cooling fan.

- Turn on the power of the detector.